Placing the Rotor on the Centrifuge

Warning: If the pins of the rotor are not properly aligned on the centrifuge before running, damage can result!

- Turn rotor upside down and note the location/orientation of its pins.

- The pins on the rotor should be positioned parallel to the pins of the centrifuge drive shaft. (i.e The pins should touch at both points and not be in-between the drive shaft pins-See Drawing Below)

Black lines/dots  represent the drive shaft, red lines/dots are  rotor pins

Correct Alignment					Improper Alignment

- Lower the rotor carefully on the Drive shaft, with its pins in the correct orientation.

- Once lowered, check to make sure the rotor pins are hugging the drive shaft pins before tightening the rotor and running the centrifuge.

Richard Kanody

Beckman coulter, Inc

(210)422-2398

rkanody@beckman.com